## Workload
Room Type | Total Workload per Patient (Wnorm) (mA min patient-1) | Typical Number of Patients (N) (per 40 h week) |  | Total Workload per Week (Wtot) (mA min week–1)
|  | Average | Busy | Average | Busy
Rad Room (chest bucky) | 0.6 | 120 | 160 | 75 | 100
Rad Room (floor or other barriers) | 1.9 | 120 | 160 | 240 | 320
Chest Room | 0.22 | 200 | 400 | 50 | 100
Fluoroscopy Tube (R&F room) | 13 | 20 | 30 | 260 | 400
Rad Tube (R&F room) | 1.5 | 25 | 40 | 40 | 60
Mammography Room | 6.7 | 80 | 160 | 550 | 1075
Cardiac Angiography | 160 | 20 | 30 | 3.2 | 4800
Peripheral Angiography | 64 | 20 | 30 | 1.3 | 2000
Rad Room (all barriers) | 25

## W kVp
kVp | Rad Room (all barriers) | Rad Roon (Chest bucky) | Rad Room (floor or other barriers) | Fluoro. Tube (R&F room) | Rad Tube (R&F room) | Chest Room | Mammo Room | Cardiac Angiography | Peripheral Angiography
25 | 0 | 0 | 0 | 0 | 0 | 0 | 0.925 | 0 | 0
30 | 0 | 0 | 0 | 0 | 0 | 0 | 4.67 | 0 | 0
35 | 0 | 0 | 0 | 0 | 0 | 0 | 1.1 | 0 | 0
40 | 0.000138 | 0 | 0.00038 | 0 | 0 | 0 | 0 | 0 | 0
45 | 0.00071 | 0 | 0.00071 | 0 | 0.000578 | 0 | 0 | 0 | 0
50 | 0.00848 | 0.00678 | 0.0017 | 0 | 0.000765 | 0 | 0 | 0.34 | 0.0894
55 | 0.0109 | 0.000456 | 0.0104 | 0.0702 | 0.000726 | 0 | 0 | 0.42 | 0.0398
60 | 0.0981 | 0.00896 | 0.0891 | 0.113 | 0.0152 | 0 | 0 | 1.96 | 0.699
65 | 0.104 | 0.0342 | 0.07 | 0.187 | 0.0252 | 0 | 0 | 4.55 | 15
70 | 0.458 | 0.0725 | 0.385 | 0.145 | 0.0889 | 0.0202 | 0 | 6.03 | 12.2
75 | 0.501 | 0.0953 | 0.405 | 0.194 | 0.224 | 0.00236 | 0 | 8.02 | 15.3
80 | 0.56 | 0.14 | 0.42 | 1.72 | 0.428 | 0 | 0 | 25.4 | 11
85 | 0.315 | 0.0662 | 0.249 | 2.19 | 0.218 | 0.000783 | 0 | 40.3 | 4.09
90 | 0.176 | 0.0141 | 0.162 | 1.46 | 0.0533 | 0 | 0 | 21 | 3.43
95 | 0.0218 | 0.00351 | 0.0182 | 1.15 | 0.0489 | 0 | 0 | 10.6 | 0.673
100 | 0.0155 | 0.000884 | 0.0146 | 1.12 | 0.0587 | 0.0301 | 0 | 7.4 | 1.53
105 | 0.00348 | 0.00197 | 0.00151 | 0.964 | 0.0105 | 0 | 0 | 7.02 | 0.0927
110 | 0.0105 | 0.00991 | 0.000551 | 0.747 | 0.0646 | 0.0214 | 0 | 6.59 | 0.0305
115 | 0.041 | 0.0374 | 0.00369 | 1.44 | 0.0029 | 0.0936 | 0 | 13.8 | 0
120 | 0.0699 | 0.0512 | 0.0187 | 0.937 | 0.104 | 0.0474 | 0 | 3.35 | 0
125 | 0.0484 | 0.0481 | 0.000347 | 0.138 | 0.0813 | 0 | 0 | 2.75 | 0
130 | 0.00184 | 0.00171 | 0.000125 | 0.153 | 0.0446 | 0 | 0 | 0.031 | 0
135 | 0.00773 | 0.00773 | 0 | 0.145 | 0.00947 | 0 | 0 | 0 | 0
140 | 0 | 0 | 0 | 0.0192 | 0.00426 | 0 | 0 | 0 | 0
Total Workload | 25 | 0.6 | 1.9 | 13 | 1.5 | 0.22 | 6.7 | 160 | 64
Patietns per week | 110 |  |  | 18 | 23 | 210 | 47 | 19 | 21

## Uns Air Kerma
Workload Distribution | Wnorm (mA min patient–1) | F (cm2) | df(m) | Leakage | Side-Scatter | Leakage and Side-Scatter (Ksec) | Forward/ Backscatter | Leakage and Forward/ Backscatter (Ksec)
Rad Room (all barriers) | 2.5 | 1000 | 1 | 0.00053 | 0.034 | 0.034 | 0.048 | 0.049
Rad Room (chest bucky) | 0.6 | 1535 | 1.83 | 0.00039 | 0.0049 | 0.0053 | 0.0069 | 0.0073
Rad Room (floor or other barriers) | 1.9 | 1000 | 1 | 0.00014 | 0.023 | 0.023 | 0.033 | 0.033
Fluoroscopy Tube (R&F room) | 13 | 730 | 0.8 | 0.012 | 0.31 | 0.32 | 0.44 | 0.46
Rad Tube (R&F room) | 1.5 | 1000 | 1 | 0.00094 | 0.028 | 0.029 | 0.039 | 0.04
Chest Room | 0.22 | 1535 | 20 | 0.00038 | 0.0023 | 0.0027 | 0.0032 | 0.0036
Mammography Room | 6.7 | 720 | 0.58 | 1.1e-05 | 0.011 | 0.011 | 0.042 | 0.049
Cardiac Angiography | 160 | 730 | 0.9 | 0.088 | 2.6 | 2.7 | 3.7 | 3.8
Peripheral Angiography | 64 | 730 | 0.9 | 0.0034 | 0.66 | 0.66 | 0.95 | 0.95

## unshielding prim
Workload Distribution | Wnorm (mA min/patient) | Kp (mGy/patient)
Rad Room (chest bucky) | 0.6 | 2.3
Rad Room (floor or other barriers) | 1.9 | 5.2
Rad Tube (R&F Room) | 1.5 | 5.9
Chest Room | 0.22 | 1.2

## sec abc
| Lead |  |  | Concrete |  |  | Gypsum Wallboard |  |  | Steel |  |  | Plate Glass |  |  | Wood
Workload Distribution | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ
30 | 38.79 | 180 | 0.356 | 0.3174 | 1.725 | 0.3705 | 0.1198 | 0.7137 | 0.3703 | 7.408 | 42.49 | 0.4061 | 0.306 | 1.62 | 0.3793 | 0.02159 | 0.03971 | 0.2852
50 | 8.801 | 27.28 | 0.2957 | 0.0903 | 0.1712 | 0.2324 | 0.0388 | 0.0873 | 0.5105 | 1.817 | 4.84 | 0.4021 | 0.09721 | 0.1799 | 0.4912 | 0.01076 | 0.001862 | 1.17
70 | 5.369 | 23.49 | 0.5883 | 0.059 | 0.1697 | 0.3849 | 0.023 | 0.0716 | 0.73 | 0.7149 | 3.798 | 0.5381 | 0.05791 | 0.1357 | 0.5968 | 0.00855 | 0.000539 | 1.194
100 | 2.507 | 15.33 | 0.9124 | 0.0395 | 0.0844 | 0.5191 | 0.0147 | 0.04 | 0.9752 | 0.3424 | 2.456 | 0.9388 | 0.04279 | 0.08948 | 1.029 | 0.00723 | 0.000894 | 1.316
125 | 2.233 | 7.888 | 0.7295 | 0.0351 | 0.066 | 0.7832 | 0.012 | 0.0267 | 1.079 | 0.2138 | 1.69 | 1.086 | 0.03654 | 0.0579 | 1.093 | 0.006587 | -0.00114 | 1.172
150 | 1.791 | 5.478 | 0.5678 | 0.0324 | 0.0775 | 1.566 | 0.0104 | 0.0202 | 1.135 | 0.1511 | 1.124 | 1.151 | 0.03267 | 0.04074 | 1.134 | 0.006027 | -0.00163 | 1.44
Rad Room (all barriers) | 2.298 | 17.38 | 0.6193 | 0.0361 | 0.1433 | 0.56 | 0.0138 | 0.057 | 0.7937 | 0.2191 | 3.49 | 0.7358 | 0.03873 | 0.1054 | 0.6397 | 0.007552 | 0.000737 | 1.044
Rad Room (chest bucky) | 2.256 | 13.8 | 0.8837 | 0.0356 | 0.1079 | 0.7705 | 0.0127 | 0.0445 | 1.049 | 0.2211 | 2.836 | 1.123 | 0.03749 | 0.0871 | 0.9086 | 0.007058 | 0.000229 | 1.875
Rad Room (floor or other barriers) | 2.513 | 17.34 | 0.4994 | 0.0392 | 0.1464 | 0.4486 | 0.0164 | 0.0608 | 0.7472 | 0.244 | 3.012 | 0.5019 | 0.04299 | 0.107 | 0.5538 | 0.007887 | 0.000877 | 0.98
Fluoroscopy Tube (R&F room) | 2.322 | 12.91 | 0.7575 | 0.0363 | 0.0936 | 0.5955 | 0.0133 | 0.041 | 0.9566 | 0.2331 | 2.213 | 0.8051 | 0.03886 | 0.08091 | 0.852 | 0.007057 | 0.000422 | 1.664
Rad Tube (R&F room) | 2.272 | 13.6 | 0.7184 | 0.0356 | 0.1114 | 0.662 | 0.0129 | 0.0457 | 0.9355 | 0.2149 | 2.695 | 0.8768 | 0.03762 | 0.08857 | 0.8087 | 0.007102 | 0.000345 | 1.698
Chest Room | 2.288 | 9.848 | 1.054 | 0.0364 | 0.0659 | 0.7543 | 0.013 | 0.0297 | 1.195 | 0.2518 | 1.829 | 1.273 | 0.03866 | 0.0627 | 1.128 | 0.007485 | -0.00081 | 0.09459
Mammography Room | 29.91 | 184.4 | 0.355 | 0.2539 | 1.8411 | 0.3924 | 0.0883 | 0.7526 | 0.3786 | 5.798 | 44.12 | 0.4124 | 0.2404 | 1.709 | 0.3918 | 0.01888 | 0.04172 | 0.2903
Cardiac Angiography | 2.354 | 14.94 | 0.7481 | 0.0371 | 0.1067 | 0.5733 | 0.0139 | 0.0464 | 0.9185 | 0.253 | 2.592 | 0.7999 | 0.04001 | 0.0903 | 0.8019 | 0.007266 | 0.000674 | 1.235
Peripheral Angiography | 2.661 | 19.54 | 0.5094 | 0.04219 | 0.1559 | 0.4472 | 0.01747 | 0.06422 | 0.7299 | 0.3579 | 3.466 | 0.56 | 0.04612 | 0.1198 | 0.5907 | 0.008079 | 0.000847 | 0.9742

## prim abc
| Lead |  |  | Concrete |  |  | Gypsum Wallboard |  |  | Steel |  |  | Plate Glass |  |  | Wood
Workload Distribution | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ
25 | 49.52 | 194 | 0.3037 | 0.3904 | 1.645 | 0.2757 | 0.1576 | 0.7175 | 0.3048 | 9.364 | 42.25 | 0.3202 | 0.3804 | 1.543 | 0.2869 | 0.0223 | 0.0434 | 0.1937
30 | 38.8 | 178 | 0.3473 | 0.3173 | 1.698 | 0.3593 | 0.1208 | 0.7043 | 0.3613 | 7.406 | 41.93 | 0.3959 | 0.3061 | 1.599 | 0.3693 | 0.02166 | 0.03966 | 0.2843
35 | 29.55 | 164.7 | 0.3948 | 0.2528 | 1.807 | 0.4648 | 0.0878 | 0.6988 | 0.4245 | 5.716 | 43.41 | 0.4857 | 0.2396 | 1.694 | 0.4683 | 0.01901 | 0.03873 | 0.3732
40 | 0 | 0 | 0 | 0.1297 | 0.178 | 0.2189 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
45 | 0 | 0 | 0 | 0.1095 | 0.1741 | 0.2269 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
50 | 8.801 | 27.28 | 0.2957 | 0.09032 | 0.1712 | 0.2324 | 0.03883 | 0.0873 | 0.5105 | 1.817 | 4.84 | 0.4021 | 0.09721 | 0.1799 | 0.4912 | 0.01076 | 0.001862 | 1.17
55 | 7.839 | 25.92 | 0.3499 | 0.07422 | 0.1697 | 0.2454 | 0.03419 | 0.08315 | 0.5606 | 1.493 | 4.515 | 0.4293 | 0.08552 | 0.1661 | 0.5112 | 0.01012 | 0.001404 | 1.269
60 | 6.951 | 24.89 | 0.4198 | 0.06251 | 0.1692 | 0.2733 | 0.02985 | 0.07961 | 0.6169 | 1.183 | 4.219 | 0.4571 | 0.07452 | 0.1539 | 0.5304 | 0.009512 | 0.0009672 | 1.333
65 | 6.13 | 24.09 | 0.5019 | 0.05528 | 0.1696 | 0.3217 | 0.02609 | 0.07597 | 0.6756 | 0.9172 | 3.982 | 0.4922 | 0.06514 | 0.1443 | 0.5582 | 0.00899 | 0.000647 | 1.353
70 | 5.369 | 23.49 | 0.5881 | 0.05087 | 0.1696 | 0.3847 | 0.02302 | 0.07163 | 0.7299 | 0.7149 | 3.798 | 0.5378 | 0.05791 | 0.1357 | 0.5967 | 0.00855 | 0.000539 | 1.194
75 | 4.666 | 26.69 | 0.6618 | 0.04797 | 0.1663 | 0.4492 | 0.02066 | 0.06649 | 0.775 | 0.5793 | 3.629 | 0.5908 | 0.05291 | 0.128 | 0.6478 | 0.008203 | 0.0006421 | 1.062
80 | 4.04 | 21.69 | 0.7187 | 0.04583 | 0.1549 | 0.4926 | 0.01886 | 0.06093 | 0.8103 | 0.4921 | 3.428 | 0.6427 | 0.04955 | 0.1208 | 0.7097 | 0.007903 | 0.000864 | 0.9703
85 | 3.504 | 20.37 | 0.755 | 0.04398 | 0.1348 | 0.4943 | 0.01746 | 0.05558 | 0.8392 | 0.4355 | 3.178 | 0.6861 | 0.04721 | 0.114 | 0.7786 | 0.007686 | 0.001056 | 1.015
90 | 3.067 | 18.83 | 0.7726 | 0.04228 | 0.1137 | 0.469 | 0.01633 | 0.05039 | 0.8585 | 0.3971 | 2.913 | 0.7204 | 0.0455 | 0.1077 | 0.8522 | 0.007511 | 0.001159 | 1.081
95 | 2.731 | 17.07 | 0.7714 | 0.04068 | 0.09705 | 0.4406 | 0.01543 | 0.04571 | 0.8763 | 0.3681 | 2.654 | 0.7461 | 0.0441 | 0.1013 | 0.9222 | 0.007345 | 0.001133 | 1.116
100 | 2.5 | 15.28 | 0.7557 | 0.03925 | 0.08567 | 0.4273 | 0.01466 | 0.04171 | 0.8939 | 0.3415 | 2.42 | 0.7645 | 0.04278 | 0.09466 | 0.9791 | 0.00723 | 0.0009343 | 1.309
105 | 2.364 | 13.41 | 0.7239 | 0.03808 | 0.07862 | 0.4394 | 0.01397 | 0.03815 | 0.908 | 0.3135 | 2.227 | 0.7788 | 0.04143 | 0.08751 | 1.014 | 0.00705 | 0.0006199 | 1.365
110 | 2.296 | 11.7 | 0.6827 | 0.03715 | 0.07436 | 0.4752 | 0.01336 | 0.03521 | 0.9244 | 0.2849 | 2.061 | 0.7897 | 0.04008 | 0.08047 | 1.03 | 0.006921 | 0.0001976 | 3.309
115 | 2.265 | 10.21 | 0.6363 | 0.03636 | 0.07201 | 0.5319 | 0.01283 | 0.03271 | 0.9423 | 0.2579 | 1.922 | 0.8008 | 0.03878 | 0.07394 | 1.033 | 0.006864 | -0.0003908 | 0.6469
120 | 2.246 | 8.95 | 0.5873 | 0.03566 | 0.07109 | 0.6073 | 0.01235 | 0.03047 | 0.9566 | 0.2336 | 1.797 | 0.8116 | 0.03758 | 0.06808 | 1.031 | 0.006726 | -0.0008308 | 1.006
125 | 2.219 | 7.923 | 0.5386 | 0.03502 | 0.07113 | 0.6974 | 0.01192 | 0.02863 | 0.9684 | 0.213 | 1.677 | 0.8247 | 0.03652 | 0.06304 | 1.031 | 0.006584 | -0.0001214 | 1.192
130 | 2.17 | 7.094 | 0.4909 | 0.03445 | 0.0716 | 0.7969 | 0.01155 | 0.02702 | 0.9802 | 0.1969 | 1.557 | 0.8309 | 0.03561 | 0.05874 | 1.037 | 0.006472 | -0.0001539 | 1.285
135 | 2.102 | 6.45 | 0.4469 | 0.03394 | 0.07263 | 0.9099 | 0.01122 | 0.02561 | 0.9901 | 0.1838 | 1.44 | 0.8391 | 0.03481 | 0.05519 | 1.049 | 0.006306 | -0.0001731 | 1.465
140 | 2.009 | 5.916 | 0.4018 | 0.03345 | 0.07476 | 1.047 | 0.01088 | 0.02436 | 0.9964 | 0.1724 | 1.328 | 0.8458 | 0.03407 | 0.05145 | 1.057 | 0.006191 | -0.0001849 | 1.53
145 | 1.895 | 5.498 | 0.358 | 0.03296 | 0.07875 | 1.224 | 0.01056 | 0.02313 | 0.9987 | 0.1616 | 1.225 | 0.8519 | 0.03336 | 0.04795 | 1.063 | 0.006115 | -0.0001869 | 1.498
150 | 1.757 | 5.177 | 0.3156 | 0.03243 | 0.08599 | 1.467 | 0.0103 | 0.02198 | 1.013 | 0.1501 | 1.132 | 0.8566 | 0.03266 | 0.04491 | 1.073 | 0.00602 | -0.0001752 | 1.483
Rad Room (all barriers) | 2.346 | 15.9 | 0.4982 | 0.03626 | 0.1429 | 0.4932 | 0.0142 | 0.05781 | 0.7445 | 0.2163 | 3.101 | 0.5745 | 0.03907 | 0.1069 | 0.594 | 0.007616 | 0.000767 | 1.027
Rad Room (chest bucky) | 2.264 | 13.08 | 0.56 | 0.03552 | 0.1177 | 0.6007 | 0.01278 | 0.04848 | 0.8609 | 0.2179 | 2.677 | 0.7209 | 0.03762 | 0.09751 | 0.7867 | 0.007142 | 0.000308 | 1.617
Rad Room (floor or other barriers) | 2.651 | 16.56 | 0.4585 | 0.03994 | 0.1448 | 0.4231 | 0.01679 | 0.06124 | 0.7356 | 0.2535 | 2.74 | 0.4297 | 0.04361 | 0.1082 | 0.5463 | 0.007915 | 0.00088 | 0.979
Fluoroscopy Tube (R&F room) | 2.347 | 12.67 | 0.6149 | 0.03616 | 0.09721 | 0.5186 | 0.0134 | 0.04283 | 0.8796 | 0.2323 | 2.19 | 0.6509 | 0.03901 | 0.08588 | 0.8081 | 0.007089 | 0.000474 | 1.58
Rad Tube (R&F room) | 2.295 | 13 | 0.5573 | 0.03549 | 0.1164 | 0.5774 | 0.013 | 0.04778 | 0.8485 | 0.2126 | 2.568 | 0.6788 | 0.03778 | 0.09365 | 0.7483 | 0.007162 | 0.000411 | 1.541
Chest Room | 2.283 | 10.74 | 0.637 | 0.03622 | 0.07766 | 0.5404 | 0.01286 | 0.03505 | 0.9356 | 0.25 | 1.989 | 0.7721 | 0.03866 | 0.07721 | 0.9843 | 0.00765 | -0.00098 | 0.08083
Mammography Room | 30.6 | 17.76 | 0.3308 | 0.2577 | 1.765 | 0.3644 | 0.09148 | 0.709 | 0.3459 | 5.998 | 42.91 | 0.3927 | 0.2467 | 1.654 | 0.3694 | 0.001914 | 0.04166 | 0.2858
Cardiac Angiography | 2.389 | 14.26 | 0.5948 | 0.03717 | 0.1087 | 0.4879 | 0.01409 | 0.04814 | 0.8419 | 0.2533 | 2.461 | 0.6243 | 0.04025 | 0.09482 | 0.7523 | 0.007303 | 0.000722 | 1.204
Peripheral Angiography | 2.728 | 18.52 | 0.4614 | 0.04292 | 0.1538 | 0.4236 | 0.01774 | 0.06449 | 0.7158 | 0.367 | 3.26 | 0.5036 | 0.04642 | 0.1203 | 0.5763 | 0.008103 | 0.000844 | 0.9754

## CT abc
| Lead |  |  | Concrete
Workload Distribution | α (mm-1) | β (mm-1) | γ | α (mm-1) | β (mm-1) | γ
120 | 2.246 | 5.73 | 0.547 | 0.0383 | 0.0142 | 0.658
140 | 2.009 | 3.99 | 0.342 | 0.0336 | 0.0122 | 0.519

## Occupancy Factor ( T )
Location | Occupancy Factor ( T )
Administrative or clerical offices | 1
Laboratories | 1
Pharmacies and other work areas fully occupied by an individual | 1
Receptionist areas | 1
Attended waiting rooms | 1
Children’s indoor play areas | 1
Adjacent x-ray rooms | 1
Film reading areas | 1
Nurse’s stations | 1
X-ray control rooms | 1
Rooms used for patient examinations and treatments | 0.5
Corridors | 0.2
Patient rooms | 0.2
Employee lounges | 0.2
Staff rest rooms | 0.2
Corridor doors | 0.13
Public toilets | 0.05
Unattended vending areas | 0.05
Storage rooms | 0.05
Outdoor areas with seating | 0.05
Unattended waiting rooms | 0.05
Patient holding areas | 0.05
Outdoor areas with only transient pedestrian or vehicular traffic | 0.03
Unattended parking lots | 0.03
Vehicular drop off areas (unattended) | 0.03
Attics | 0.03
Stairways | 0.03
Unattended elevators | 0.03
Janitor’s closets | 0.03

## Primary beam use factors (U) fo
Barrier | Use Factor (U) | Apply to workload Distribution
Floor | 0.89 | Rad Room (floor or other barriers)
Cross-table wall | 0.09 | Rad Room (floor or other barriers)
Wall No 3 | 0.02 | Rad Room (floor or other barriers)
Chest Image receptor | 1 | Rad Room (chest bucky)

## Equiv. thickness of prim pres
Application | xpre (in mm)
| Lead | Concrete | Steel
Image receptor in radiographic table or wall-mounted cassette holder (attenuation by grid, cassette, and image-receptor supporting structures) | 0.85 | 72 | 7
Cross-table lateral (attenuation by grid and cassette only) | 0.3 | 30 | 2

## Sheet1
x | =

## OF BIR
Control rooms | 1
Reception areas | 1
Nurses stations | 1
Offices | 1
shops | 1
living quarters | 1
children's indoor play areas | 1
occupied space in nearby building | 1
Staff rooms
Adjacent Wards
Clinic rooms
Reporting areas
Corridors
Storerooms
Stairways
Changing rooms
Unattended car parks
Unattended waiting rooms
Toilets
bathrooms